## Data 1
| Survey No | INIS-121019-783 |  | Item Surveyed |  | Characterization of Fume Hoods:1661, 1669, 1665, 1817, and 2055
| Date | 2019-12-10 |  | Comments |  | This surey includes 5 fume hoods out of a total of 41. See attached map for the survey description.
| Survey Tech | J.Walther, A.Peden
| Count Room Tech | P. Ray |  | Parameters |  | Gamma |  |  | Total Activity |  |  |  | Removable Activity |  |  |  | 1.4142
| Date Counted | 2019-12-10 |  |  |  | CPM | µR/hr | µRem/hr | Alpha |  | Beta-Gamma |  | Alpha |  | Beta-Gamma |  | 0
| Survey Type | Characterization |  |  | Instrument Model | 2221 |  |  | 2360/43-93 |  | 2360/43-93 |  | 2929/43-10-1 |  | 2929/43-10-1 |  | 0
| Level of Posting | None |  |  | Instrument SN | 190201 |  |  | 234860/PR289402 |  | 234860/PR289402 |  | 143878/PR147628 |  | 143878/PR147628 |  | 1.4142
Notes |  |  | Total Activity |  | Cal. Due Date | 2020-06-14 |  |  | 2020-06-18 |  | 2020-06-18 |  | 2020-11-08 |  | 2020-11-08 |  | 18.2757
PCF = Probe Correction Factor |  |  | dpm = (cpm - Bcpm)/(eff * PCF) |  | Efficiency |  |  |  | 20.93% |  | 33.78% |  | 37.87% |  | 42.14% |  | 25.8457
Tb = Background count time |  |  | Removable Activity |  | Background Counts | 3318 |  |  | 0 |  | 334 |  | 6 |  | 2566 |  | 1.0083
Ts = Sample count time |  |  | dpm = (cpm - Bcpm)/ eff |  | PCF |  |  |  | 1 |  | 1 |  | 1 |  | 1
Rb = Background count rate |  |  | *dpm results are per 100cm2 |  | Tb |  |  |  | 1 |  | 1 |  | 60 |  | 60 |  | #VALUE!
Bcpm = Background cpm |  |  |  |  | Ts |  |  |  | 1 |  | 1 |  | 1 |  | 1 |  | 1.0083
MDCR = Minimum Detectable Count Rate (net cpm) |  |  |  | MDCR |  |  |  |  | 3 |  | 88.0323 |  | 4.049 |  | 24.6939 |  | #REF!
MDC = Minimum Detectable Concentration (dpm per 100cm2) |  |  |  | MDC |  |  |  |  | 14.3335 |  | 260.6049 |  | 10.6919 |  | 58.5997
No. | Descriptions |  |  |  |  | cpm | µR/hr | µRem/hr | gross counts | *dpm | gross counts | *dpm | gross counts | *dpm | gross counts | *dpm
1 | Rm 327 (UW FH#1661) Inside of Duct (A) |  |  |  |  | 2860 |  |  | 1 | 4.7778 | 285 | -145.0562 | 0 | -0.26406 | 49 | 14.792
2 | Back (B) | [11 empty] | 0 | -0.26406 | 37 | -13.6845
3 | Back (C) | [11 empty] | 2 | 5.0172 | 44 | 2.9268
Background 2221: 3318 cpm, 2360: 0/334 cpm
2221 Scan Range: 2403 - 2989 cpm

## Data 2
| Survey No | INIS-121019-783 |  | Item Surveyed |  | Characterization of Fume Hoods:1661, 1669, 1665, 1817, and 2055
| Date | 2019-12-10 |  | Comments |  | This surey includes 5 fume hoods out of a total of 41. See attached map for the survey description.
| Survey Tech | J.Walther, A.Peden
| Count Room Tech | P. Ray |  | Parameters |  | Gamma |  |  | Total Activity |  |  |  | Removable Activity |  |  |  | 1.4142
| Date Counted | 2019-12-10 |  |  |  | CPM | µR/hr | µRem/hr | Alpha |  | Beta-Gamma |  | Alpha |  | Beta-Gamma |  | 1.4142
| Survey Type | Characterization |  |  | Instrument Model | 2221 |  |  | 2360/43-93 |  | 2360/43-93 |  | 2929/43-10-1 |  | 2929/43-10-1 |  | 2
| Level of Posting | None |  |  | Instrument SN | 190201 |  |  | 234860/PR289402 |  | 234860/PR289402 |  | 143878/PR147628 |  | 143878/PR147628 |  | 1.4142
Notes |  |  | Total Activity |  | Cal. Due Date | 2020-06-14 |  |  | 2020-06-18 |  | 2020-06-18 |  | 2020-11-08 |  | 2020-11-08 |  | 18.4391
PCF = Probe Correction Factor |  |  | dpm = (cpm - Bcpm)/(eff * PCF) |  | Efficiency |  |  |  | 20.93% |  | 33.78% |  | 37.87% |  | 42.14% |  | 26.0768
Tb = Background count time |  |  | Removable Activity |  | Background Counts | 3972 |  |  | 2 |  | 340 |  | 6 |  | 2566 |  | 1.0083
Ts = Sample count time |  |  | dpm = (cpm - Bcpm)/ eff |  | PCF |  |  |  | 1 |  | 1 |  | 1 |  | 1
Rb = Background count rate |  |  | *dpm results are per 100cm2 |  | Tb |  |  |  | 1 |  | 1 |  | 60 |  | 60 |  | #VALUE!
Bcpm = Background cpm |  |  |  |  | Ts |  |  |  | 1 |  | 1 |  | 1 |  | 1 |  | 1.0083
MDCR = Minimum Detectable Count Rate (net cpm) |  |  |  | MDCR |  |  |  |  | 9.58 |  | 88.7927 |  | 4.049 |  | 24.6939 |  | #REF!
MDC = Minimum Detectable Concentration (dpm per 100cm2) |  |  |  | MDC |  |  |  |  | 45.7716 |  | 262.8558 |  | 10.6919 |  | 58.5997
No. | Descriptions |  |  |  |  | cpm | µR/hr | µRem/hr | gross counts | *dpm | gross counts | *dpm | gross counts | *dpm | gross counts | *dpm
4 | Rm SB-049 (UW FH#1669) Inside of Duct (A) |  |  |  |  | 3862 |  |  | 1 | -4.7778 | 298 | -124.3339 | 0 | -0.26406 | 35 | -18.4306
5 | Back (B) | [11 empty] | 0 | -0.26406 | 32 | -25.5498
6 | Back (C) | [11 empty] | 0 | -0.26406 | 36 | -16.0576
Background 2221: 3972 cpm, 2360: 2/340 cpm
2221 Scan Range: 3408 - 3946 cpm

## Data 3
| Survey No | INIS-121019-783 |  | Item Surveyed |  | Characterization of Fume Hoods:1661, 1669, 1665, 1817, and 2055
| Date | 2019-12-10 |  | Comments |  | This surey includes 5 fume hoods out of a total of 41. See attached map for the survey description.
| Survey Tech | J.Walther, A.Peden
| Count Room Tech | P. Ray |  | Parameters |  | Gamma |  |  | Total Activity |  |  |  | Removable Activity |  |  |  | 1.4142
| Date Counted | 2019-12-10 |  |  |  | CPM | µR/hr | µRem/hr | Alpha |  | Beta-Gamma |  | Alpha |  | Beta-Gamma |  | 1
| Survey Type | Characterization |  |  | Instrument Model | 2221 |  |  | 2360/43-93 |  | 2360/43-93 |  | 2929/43-10-1 |  | 2929/43-10-1 |  | 1.4142
| Level of Posting | None |  |  | Instrument SN | 190201 |  |  | 234860/PR289402 |  | 234860/PR289402 |  | 143878/PR147628 |  | 143878/PR147628 |  | 1.4142
Notes |  |  | Total Activity |  | Cal. Due Date | 2020-06-14 |  |  | 2020-06-18 |  | 2020-06-18 |  | 2020-11-08 |  | 2020-11-08 |  | 17.2627
PCF = Probe Correction Factor |  |  | dpm = (cpm - Bcpm)/(eff * PCF) |  | Efficiency |  |  |  | 20.93% |  | 33.78% |  | 37.87% |  | 42.14% |  | 24.4131
Tb = Background count time |  |  | Removable Activity |  | Background Counts | 3451 |  |  | 1 |  | 298 |  | 6 |  | 2566 |  | 1.0083
Ts = Sample count time |  |  | dpm = (cpm - Bcpm)/ eff |  | PCF |  |  |  | 1 |  | 1 |  | 1 |  | 1
Rb = Background count rate |  |  | *dpm results are per 100cm2 |  | Tb |  |  |  | 1 |  | 1 |  | 60 |  | 60 |  | #VALUE!
Bcpm = Background cpm |  |  |  |  | Ts |  |  |  | 1 |  | 1 |  | 1 |  | 1 |  | 1.0083
MDCR = Minimum Detectable Count Rate (net cpm) |  |  |  | MDCR |  |  |  |  | 7.6528 |  | 83.3191 |  | 4.049 |  | 24.6939 |  | #REF!
MDC = Minimum Detectable Concentration (dpm per 100cm2) |  |  |  | MDC |  |  |  |  | 36.5636 |  | 246.6523 |  | 10.6919 |  | 58.5997
No. | Descriptions |  |  |  |  | cpm | µR/hr | µRem/hr | gross counts | *dpm | gross counts | *dpm | gross counts | *dpm | gross counts | *dpm
7 | Rm B-004 (UW FH#1665) Inside of Duct (A) |  |  |  |  | 3348 |  |  | 2 | 4.7778 | 288 | -29.6033 | 0 | -0.26406 | 53 | 24.2841
8 | Back (B) | [11 empty] | 0 | -0.26406 | 36 | -16.0576
9 | Back (C) | [11 empty] | 0 | -0.26406 | 42 | -1.8193
Background 2221: 3451 cpm, 2360: 1/298 cpm
2221 Scan Range: 2579 - 3064 cpm

## Data 4
| Survey No | INIS-121019-783 |  | Item Surveyed |  | Characterization of Fume Hoods:1661, 1669, 1665, 1817, and 2055
| Date | 2019-12-10 |  | Comments |  | This surey includes 5 fume hoods out of a total of 41. See attached map for the survey description.
| Survey Tech | J.Walther, A.Peden
| Count Room Tech | P. Ray |  | Parameters |  | Gamma |  |  | Total Activity |  |  |  | Removable Activity |  |  |  | 1.4142
| Date Counted | 2019-12-10 |  |  |  | CPM | µR/hr | µRem/hr | Alpha |  | Beta-Gamma |  | Alpha |  | Beta-Gamma |  | 0
| Survey Type | Characterization |  |  | Instrument Model | 2221 |  |  | 2360/43-93 |  | 2360/43-93 |  | 2929/43-10-1 |  | 2929/43-10-1 |  | 0
| Level of Posting | None |  |  | Instrument SN | 190201 |  |  | 234860/PR289402 |  | 234860/PR289402 |  | 143878/PR147628 |  | 143878/PR147628 |  | 1.4142
Notes |  |  | Total Activity |  | Cal. Due Date | 2020-06-14 |  |  | 2020-06-18 |  | 2020-06-18 |  | 2020-11-08 |  | 2020-11-08 |  | 17.7482
PCF = Probe Correction Factor |  |  | dpm = (cpm - Bcpm)/(eff * PCF) |  | Efficiency |  |  |  | 20.93% |  | 33.78% |  | 37.87% |  | 42.14% |  | 25.0998
Tb = Background count time |  |  | Removable Activity |  | Background Counts | 3036 |  |  | 0 |  | 315 |  | 6 |  | 2566 |  | 1.0083
Ts = Sample count time |  |  | dpm = (cpm - Bcpm)/ eff |  | PCF |  |  |  | 1 |  | 1 |  | 1 |  | 1
Rb = Background count rate |  |  | *dpm results are per 100cm2 |  | Tb |  |  |  | 1 |  | 1 |  | 60 |  | 60 |  | #VALUE!
Bcpm = Background cpm |  |  |  |  | Ts |  |  |  | 1 |  | 1 |  | 1 |  | 1 |  | 1.0083
MDCR = Minimum Detectable Count Rate (net cpm) |  |  |  | MDCR |  |  |  |  | 3 |  | 85.5783 |  | 4.049 |  | 24.6939 |  | #REF!
MDC = Minimum Detectable Concentration (dpm per 100cm2) |  |  |  | MDC |  |  |  |  | 14.3335 |  | 253.3403 |  | 10.6919 |  | 58.5997
No. | Descriptions |  |  |  |  | cpm | µR/hr | µRem/hr | gross counts | *dpm | gross counts | *dpm | gross counts | *dpm | gross counts | *dpm
10 | Rm B-018C (UW FH#1817) Inside of Duct (A) |  |  |  |  | 2676 |  |  | 0 | 0 | 271 | -130.2546 | 0 | -0.26406 | 41 | -4.1924
11 | Back (B) | [11 empty] | 0 | -0.26406 | 38 | -11.3115
12 | Back (C) | [11 empty] | 0 | -0.26406 | 49 | 14.792
Background 2221: 3036 cpm, 2360: 0/315 cpm
2221 Scan Range: 2358 - 2610 cpm

## Data 5
| Survey No | INIS-121019-783 |  | Item Surveyed |  | Characterization of Fume Hoods:1661, 1669, 1665, 1817, and 2055
| Date | 2019-12-10 |  | Comments |  | This surey includes 5 fume hoods out of a total of 41. See attached map for the survey description.
| Survey Tech | J.Walther, A.Peden
| Count Room Tech | P. Ray |  | Parameters |  | Gamma |  |  | Total Activity |  |  |  | Removable Activity |  |  |  | 1.4142
| Date Counted | 2019-12-10 |  |  |  | CPM | µR/hr | µRem/hr | Alpha |  | Beta-Gamma |  | Alpha |  | Beta-Gamma |  | 0
| Survey Type | Characterization |  |  | Instrument Model | 2221 |  |  | 2360/43-93 |  | 2360/43-93 |  | 2929/43-10-1 |  | 2929/43-10-1 |  | 0
| Level of Posting | None |  |  | Instrument SN | 190201 |  |  | 234860/PR289402 |  | 234860/PR289402 |  | 143878/PR147628 |  | 143878/PR147628 |  | 1.4142
Notes |  |  | Total Activity |  | Cal. Due Date | 2020-06-14 |  |  | 2020-06-18 |  | 2020-06-18 |  | 2020-11-08 |  | 2020-11-08 |  | 17.1172
PCF = Probe Correction Factor |  |  | dpm = (cpm - Bcpm)/(eff * PCF) |  | Efficiency |  |  |  | 20.93% |  | 33.78% |  | 37.87% |  | 42.14% |  | 24.2074
Tb = Background count time |  |  | Removable Activity |  | Background Counts | 3200 |  |  | 0 |  | 293 |  | 6 |  | 2566 |  | 1.0083
Ts = Sample count time |  |  | dpm = (cpm - Bcpm)/ eff |  | PCF |  |  |  | 1 |  | 1 |  | 1 |  | 1
Rb = Background count rate |  |  | *dpm results are per 100cm2 |  | Tb |  |  |  | 1 |  | 1 |  | 60 |  | 60 |  | #VALUE!
Bcpm = Background cpm |  |  |  |  | Ts |  |  |  | 1 |  | 1 |  | 1 |  | 1 |  | 1.0083
MDCR = Minimum Detectable Count Rate (net cpm) |  |  |  | MDCR |  |  |  |  | 3 |  | 82.6425 |  | 4.049 |  | 24.6939 |  | #REF!
MDC = Minimum Detectable Concentration (dpm per 100cm2) |  |  |  | MDC |  |  |  |  | 14.3335 |  | 244.6491 |  | 10.6919 |  | 58.5997
No. | Descriptions |  |  |  |  | cpm | µR/hr | µRem/hr | gross counts | *dpm | gross counts | *dpm | gross counts | *dpm | gross counts | *dpm
13 | Rm B-008A (UW FH#2055) Inside of Duct (A) |  |  |  |  | 2721 |  |  | 1 | 4.7778 | 283 | -29.6033 | 1 | 2.3766 | 49 | 14.792
14 | Back (B) | [11 empty] | 0 | -0.26406 | 52 | 21.9111
15 | Back (C) | [11 empty] | 0 | -0.26406 | 36 | -16.0576
Background 2221: 3200 cpm, 2360: 0/293 cpm
2221 Scan Range: 2405 - 2686 cpm

## Map
Survey No | INIS-121019-783
Date | 2019-12-10
Survey Tech | J.Walther, A.Penden
Count Room Tech | P.Ray
Date Counted | 2019-12-10
Survey Type | Characterization
Level of Posting | None
Comments
Ceiling
Side
Back
Floor
Side
Duct
Fume Hood
Inside of Fume Hood
A
B
C
A : 1, 4, 7, 10, 13
B : 2, 5, 8, 11, 14
C : 3, 6, 9, 12, 15
Gamma scan survey with 44-10 probe and beta/gamma scan survey with 43-93 probe was performed on all accessible surfaces inside of each fume hood, and on the slits outside above the door. One minute static measurements were performed inside of the duct with the both probes. Smears were obtained from inside of the duct and locations where air was likely to stagnate behind the back panel (dusty locations, B and C).
Survey Locations and Survey Numbers